ISO 10362-1 Cigarettes - Determination of water in smoke condensates - Gas-chromatographic method

ISO 10315    Cigarettes - Determination of nicotine in smoke condensates - Gas-chromatographic method

The method of internal standard is used for determination of water in extracting 2-propanol.

In accordance with ISO 10362-1 ethanol is chosen as Internal Standard (IS).

The method of internal standard is used for determination of nicotine in extracting 2-propanol.

In accordance with ISO 10362-1 anetol is chosen as Internal Standard (IS).

Warning. The main component extracting solvent 2-propanol may be successfully used as internal standard for chromatographic quantitative analysis of water and nicotine.

Determination of water (nicotine) in the extracting 2-propanol is described by the next expression

Cw(n)=Kw(n)st⋅Sw(n)Sst⋅Сst			(1)

where  Kw(n) = kw(n) /kst - are relative response of detector.

Response factors for water (nicotine) kw(n) and for substance of IS  kst  are obtained from the next equations:

kw(n)=cw(n)2−cw(n)1sw(n)2−sw(n)1				(2)

kst=cst2−cst1sst2−sst1				(3)

where  Sw(n)  and Sst – are responses for water (nicotine) and for IS, respectively, Сw(n) and Cst – are concentrations of added water (nicotine) and IS. It is necessary to stress the next fact. There is always some unknown concentration of water in the extracting 2-propanol. So, two measurements with two different additions of water should be done, at list.

Calculations of all necessary coefficients, of reproducibility, of validation and generation of result report are doing in the MS template Water in smoke condensates ISO 10362 Template - eng.xlt and Nicotine in smoke condensates ISO 10315 Template - eng.xlt

Determination of water in extracting 2-propanol. Setting up apparatus.

The GC Crystall2000M with software UniChrom

Carrier gas hydrogen	40 cc/min

Т oven 		135 С

Т inlet		250 С

Т det		250 С

Certified Mixtures of inspected water in 2-propanol. Concentration of water in blank extracting 2-propanol was 2,3 mg/2ml.

Experimental data are presented in the Table 1.

There is only one single file - UniChrom method –

Water in smoke condensates ISO 10362 Template - eng.$$$

This one file includes:

- data,

- instrument method or setting up apparatus,

- processing scenario, may by many different scenarios,

- libraries,

- calibrations,

- rich text description.

Resume. The calculated results of water concentrations in two cases, with using ethanol and 2-propanol as IS, doest not differ more than 2%.

Short note

The main component (solvent) as internal standard for chromatographic quantitative analysis of impurities

The impurity and main component concentrations were determined by the method of iterations in two stages.

First, a priori the mass fraction (in percents) Xst0 (%) of the main component selected as internal standard was set equal to 100 %.

The mass fraction (in percents) Xi (%) of the i-th component in the sample is given by

Xi=Ki⋅Si⋅XstSst	 ,			(1)

where Si and Sst are the chromatographic peak areas for the i-th component of interest and for the IS substance in the sample being analysed respectively;  Xst is the mass fraction of the IS  substance in the sample.

Numeric values of the relative mass factors Ki of detector response are obtained from the chromatographic data  for certified mixture with the known mass fractions of IS and the i-th component of interest according to

Ki=Sstcert⋅XicertSicert⋅Xstcert, 				(2)

where Sicert and  Sstcert are the chromatographic peak areas for the i-th component of interest and for the IS substance respectively, Xicert and  Xstcert are the mass fractions for the i-th component  and the IS substance respectively.

Second, the main component mass fraction is corrected by the next equation

Xst=100−i≠jXi ,			(3)

where j is the summation index corresponding to the main component. Then the impurity concentrations are calculated by equation (1) with the corrected value of Xst from eq. (3). The result expression may by written in the next form

Xi=Ki⋅100⋅(1−(i≠j)Ki⋅SiSst)⋅SiSst .   		(4)

P.S. Your notes, comments and proposals send to site www.unichrom.com