Chromatograph Validation Procedure in accordance with STB 1287-2001

Testing

At this stage the certified mixture with control substance of known mass concentration is measured.

We define the Peak Name and Peak Width at Halfheight. The last parameter is used for statistics noise characteristics.

In the 1-st layer there is the exemplary chromatogram of certified mixture.

In the 2-nd layer there is the reference chromatogram of certified mixture. It is used like library chromatogram.

It will be used for identification of late chromatograms.

In the 3-rd layer there is the measured noise.

Noise measurement duration must be more than 30 Peak Width at Halfheight.

The Relative Standard Deviation (RSD) of noise is calculated by the next equation:

Δx=1,96⋅i=1nxi−1nj=1nxj2(n−1)

where xi - is the noise value in the i-s moment of time, n - is the number of the measured noise values.

All chromatograms are measured and saved in one file Chromatograph Validation Template.$$$

Validation report is generated by template Chromatograph Validation Template.xlt

Testing report section is placed on the page Testing and Noise of book Chromatograph Validation Template.xlt

Repeatability. RSD calculation for Retention Time, Peak Height, Peak Area and MDL calculation

The template Chromatograph Validation Template.xlt automatically selects the best 10 from 12  measurements and calculates RSD for Retention Time, Peak Height, Peak Area and MDL. Selection criteria is described by the next equation:

εi2=ti−tt2+Ai−AA2+Si−SS2

where ti, Ai, Si - are the Retention Time, Peak Height and Peak Area corresponding, t, A, S - average Retention Time, Peak Height and Peak Area of control substance over all 12 measurements.

These chromatograms are measured and placed from the 4-th to 15-th layers.

RSD and MDL report section is placed on the page RSD and MDL of book Chromatograph Validation Template.xlt

Repeatability. RSD calculation for Retention Time, Peak Height, Peak Area and MDL calculation after 48 hours

After 48 hours RSD measurement procedure repeat again.

It is possible to reduce this time interval up to 8 hours after agreement with Metrology Committee authority person. 12 chromatograms of the same certificate mixture are measured and are placed  from the 16-th to 27-th layers. Reproduction report section is placed on the page Reproduction of book Chromatograph Validation Template.xlt

Certificate Report

The final Validation Report is generated and is placed on the page Validation Report of book Chromatograph Validation Template.xlt

P.S. Your notes, comments and proposals send to unichrom@unichrom.com